## Commands
Generated from the provided app.c command handlers. This workbook reflects the implementation in the shared code snippet, including behavior mismatches and known formatting quirks.
Timestamp placeholder 'TS' means HAL_GetTick() at response time. For delegated module strings (Laser/Peltier/Target/ADC/I2C/Calib), the exact enum names come from the respective *_GetStateAsString() / *_GetErrorCodeAsString() functions.
Section | CMD | Parameters | Return OK / Output | Return NOK | Description / Behavior | Example | Example return | Notes
LASER CONTROL | LASER_SET_ARM_EN | int/bool; >0 = arm, <=0 = disarm | [LASER_SET_ARM_EN][arg]->[OK][TS] | [LASER_SET_ARM_EN][arg]->[NOK][reason][TS] | Arms or disarms laser. If disarming while a process is active, stops calibration, target sequence, and laser first. Cannot arm while Peltier cooling is disabled. | LASER_SET_ARM_EN 1 | [LASER_SET_ARM_EN][1]->[OK][12523]
LASER CONTROL | LASER_GET_ARM_EN | ~ | [LASER_GET_ARM_EN]->[1¦0][TS] |  | Returns whether laser is armed. | LASER_GET_ARM_EN | [LASER_GET_ARM_EN]->[1][12523]
LASER CONTROL | LASER_SET_RED_DOT_EN | int/bool; >0 = on, <=0 = off | [LASER_SET_RED_DOT_EN][arg]->[OK][TS] | [LASER_SET_RED_DOT_EN][arg]->[NOK][reason][TS] | Turns red-dot mode on/off. If APP checkVacuum is enabled, enabling is blocked when vacuum is not acceptable. | LASER_SET_RED_DOT_EN 1 | [LASER_SET_RED_DOT_EN][1]->[OK][12523]
LASER CONTROL | LASER_GET_RED_DOT_EN | ~ | [LASER_GET_RED_DOT_EN]->[1¦0][TS] |  | Returns whether red-dot mode is active. | LASER_GET_RED_DOT_EN | [LASER_GET_RED_DOT_EN]->[0][12523]
LASER CONTROL | LASER_SET_CHANNEL_PWR | 808PWR,980PWR,1064PWR (CSV integers) | [LASER_SET_CHANNEL_PWR][p1,p2,p3]->[OK][TS] | [LASER_SET_CHANNEL_PWR][p1,p2,p3]->[NOK][reason][TS] | Sets 808/980/1064 power percentages and marks channels active when power > 0. | LASER_SET_CHANNEL_PWR 20,25,50 | [LASER_SET_CHANNEL_PWR][20,25,50]->[OK][12523] | Code validates each value as 1-100, so 0 is currently rejected here even though target entries allow 0-100.
LASER CONTROL | LASER_GET_CHANNEL_PWR | ~ | [LASER_GET_CHANNEL_PWR]->[808PWR,980PWR,1064PWR][TS] |  | Returns configured channel powers. | LASER_GET_CHANNEL_PWR | [LASER_GET_CHANNEL_PWR]->[20,25,50][12523]
LASER CONTROL | LASER_FIRE | duration_ms | [LASER_FIRE][duration]->[OK][TS] | [LASER_FIRE][duration]->[NOK][reason][TS] | Fires laser for given pulse length. Blocked unless app is RUNNING, no other process is active, and vacuum is acceptable when vacuum checking is enabled. | LASER_FIRE 20 | [LASER_FIRE][20]->[OK][12523]
LASER CONTROL | LASER_STOP | ~ | [LASER_STOP]->[OK][TS] |  | Stops laser immediately. | LASER_STOP | [LASER_STOP]->[OK][12523]
LASER CONTROL | LASER_IS_ACTIVE | ~ | [LASER_IS_ACTIVE]->[1¦0][TS] |  | Returns whether laser pulse is active. | LASER_IS_ACTIVE | [LASER_IS_ACTIVE]->[0][12523]
LASER CONTROL | LASER_GET_STATE | ~ | [LASER_GET_STATE]->[STATE_STRING][TS] |  | Returns laser controller state string. | LASER_GET_STATE | [LASER_GET_STATE]->[LASER_STATE_ARMED_IDLE][12523] | Exact state names come from Laser_GetStateAsString().
LASER CONTROL | LASER_GET_LAST_ERROR | ~ | [LASER_GET_LAST_ERROR]->[Error string][ERROR_CODE][TS] |  | Returns laser last error string and code. | LASER_GET_LAST_ERROR | [LASER_GET_LAST_ERROR]->[No error][LASER_ERROR_NONE][12523]
LASER CONTROL | LASER_CLEAR_ERROR | ~ | [LASER_CLEAR_ERROR]->[OK][TS] |  | Clears laser error. | LASER_CLEAR_ERROR | [LASER_CLEAR_ERROR]->[OK][12523]
PELTIER CONTROL | PELTIER_SET_COOLING_EN | int/bool; >0 = enable, <=0 = disable | [PELTIER_SET_COOLING_EN][arg]->[OK][TS] | [PELTIER_SET_COOLING_EN][arg]->[NOK][reason][TS] | Enables/disables Peltier cooling. Cannot modify while a process is active or app not RUNNING. Cannot disable while laser is armed. | PELTIER_SET_COOLING_EN 1 | [PELTIER_SET_COOLING_EN][1]->[OK][12523]
PELTIER CONTROL | PELTIER_GET_COOLING_EN | ~ | [PELTIER_GET_COOLING_EN]->[1¦0][TS] |  | Returns whether Peltier is enabled. | PELTIER_GET_COOLING_EN | [PELTIER_GET_COOLING_EN]->[1][12523]
PELTIER CONTROL | PELTIER_GET_STATE | ~ | [PELTIER_GET_STATE]->[STATE_STRING][TS] |  | Returns Peltier state string. | PELTIER_GET_STATE | [PELTIER_GET_STATE]->[PELTIER_STATE_ENABLED][12523]
PELTIER CONTROL | PELTIER_GET_LAST_ERROR | ~ | [PELTIER_GET_LAST_ERROR]->[Error string][ERROR_CODE][TS] |  | Returns Peltier last error string and code. | PELTIER_GET_LAST_ERROR | [PELTIER_GET_LAST_ERROR]->[No error][PELTIER_ERROR_NONE][12523]
PELTIER CONTROL | PELTIER_CLEAR_ERROR | ~ | [PELTIER_CLEAR_ERROR]->[OK][TS] |  | Clears Peltier error. | PELTIER_CLEAR_ERROR | [PELTIER_CLEAR_ERROR]->[OK][12523]
TARGET CONTROL | TARGET_SET_POS | X,Y (CSV, each 0-4095) | [TARGET_SET_POS][X,Y]->[OK][TS] | [TARGET_SET_POS][X,Y]->[NOK][reason][TS] | Sets galvo target position. Blocked while a process is active or app not RUNNING. | TARGET_SET_POS 1500,2000 | [TARGET_SET_POS][1500,2000]->[OK][12523]
TARGET CONTROL | TARGET_GET_POS | ~ | [TARGET_GET_POS]->[X,Y][TS] | [TARGET_GET_POS]->[NOK][reason][TS] | Returns commanded target position. | TARGET_GET_POS | [TARGET_GET_POS]->[1500,2000][12523]
TARGET CONTROL | TARGET_GET_FB_POS | ~ | [TARGET_GET_FB_POS]->[Gx,Gy][TS] | [TARGET_GET_FB_POS]->[NOK][Feedback position is not valid][TS] | Returns raw galvo feedback values from ADC acquisition. | TARGET_GET_FB_POS | [TARGET_GET_FB_POS]->[2033,1989][12523]
TARGET CONTROL | TARGET_SET_NEW_TARGET | X0,Y0,X1,Y1,808PWR,980PWR,1064PWR,duration_ms | [TARGET_SET_NEW_TARGET][args]->[OK][TS] | [TARGET_SET_NEW_TARGET][args]->[NOK][reason][TS] | Adds a target line/point. Power here is allowed 0-100. Pulse length must be TARGET_CTRL_MIN_PULSE_MS..maxFireDuration. | TARGET_SET_NEW_TARGET 1500,1000,2000,1500,20,10,15,50 | [TARGET_SET_NEW_TARGET][1500,1000,2000,1500,20,10,15,50]->[OK][12523]
TARGET CONTROL | TARGET_CLEAR_TARGETS | ~ | [TARGET_CLEAR_TARGETS]->[OK][TS] | [TARGET_CLEAR_TARGETS]->[NOK][reason][TS] | Clears target list. | TARGET_CLEAR_TARGETS | [TARGET_CLEAR_TARGETS]->[OK][12523]
TARGET CONTROL | TARGET_GET_TARGET | index | [TARGET_GET_TARGET]->[idx,x0,y0,x1,y1,p808,p980,p1064,duration][TS] | [TARGET_GET_TARGET][idx]->[NOK][reason][TS] | Returns target line at index. | TARGET_GET_TARGET 5 | [TARGET_GET_TARGET]->[5,1500,1000,2000,1500,20,10,15,50][12523]
TARGET CONTROL | TARGET_GET_ALL_TARGETS | ~ | Header: [TARGET_GET_ALL_TARGETS]->[Count,N][TS]; then N data lines [TARGET_GET_ALL_TARGETS]->[idx,x0,y0,x1,y1,p808,p980,p1064,duration]; then [TARGET_GET_ALL_TARGETS]->[END][TS] | [TARGET_GET_ALL_TARGETS]->[NOK][reason][TS] | Returns all valid target lines. Blocked while process is active. | TARGET_GET_ALL_TARGETS | [TARGET_GET_ALL_TARGETS]->[Count,3][12523]
TARGET CONTROL | TARGET_SET_MODE | numeric enum (TargetMode_t) | [TARGET_SET_MODE][arg]->[OK][TS] | [TARGET_SET_MODE][arg]->[NOK][reason][TS] | Sets target mode. Implementation parses an integer and casts to TargetMode_t. | TARGET_SET_MODE 0 | [TARGET_SET_MODE][0]->[OK][12523] | Important: code does NOT parse strings like SINGLE/AUTO. It expects numeric enum values.
TARGET CONTROL | TARGET_GET_MODE | ~ | [TARGET_GET_MODE]->[MODE_STRING][TS] |  | Returns target mode string. | TARGET_GET_MODE | [TARGET_GET_MODE]->[TARGET_MODE_SINGLE][12523]
TARGET CONTROL | TARGET_START | ~ | [TARGET_START]->[OK][TS] | [TARGET_START]->[NOK][reason][TS] | Starts target process. Requires RUNNING state, no active process, and acceptable vacuum when checkVacuum is enabled. | TARGET_START | [TARGET_START]->[OK][12523]
TARGET CONTROL | TARGET_HALT | ~ | [TARGET_HALT]->[OK][TS] | [TARGET_HALT]->[NOK][reason][TS] | Halts active target process. Requires app RUNNING. | TARGET_HALT | [TARGET_HALT]->[OK][12523]
TARGET CONTROL | TARGET_STOP | ~ | [TARGET_STOP]->[OK][TS] | [TARGET_STOP]->[NOK][reason][TS] | Stops target process. Requires app RUNNING. | TARGET_STOP | [TARGET_STOP]->[OK][12523]
TARGET CONTROL | TARGET_CONTINUE | ~ | [TARGET_CONTINUE]->[OK][TS] | [TARGET_CONTINUE]->[NOK][reason][TS] | Continues halted target process. Requires app RUNNING. | TARGET_CONTINUE | [TARGET_CONTINUE]->[OK][12523]
TARGET CONTROL | TARGET_GET_STATE | ~ | [TARGET_GET_STATE]->[STATE_STRING][TS] |  | Returns target controller state string. | TARGET_GET_STATE | [TARGET_GET_STATE]->[TARGET_STATE_IDLE][12523]
TARGET CONTROL | TARGET_GET_LAST_ERROR | ~ | [TARGET_GET_LAST_ERROR]->[Error string][ERROR_CODE][TS] |  | Returns target controller last error string and code. | TARGET_GET_LAST_ERROR | [TARGET_GET_LAST_ERROR]->[No error][TARGET_ERROR_NONE][12523]
TARGET CONTROL | TARGET_CLEAR_ERROR | ~ | [TARGET_CLEAR_ERROR]->[OK][TS] | [TARGET_CLEAR_ERROR]->[NOK][reason][TS] | Clears target controller error. Requires app RUNNING. | TARGET_CLEAR_ERROR | [TARGET_CLEAR_ERROR]->[OK][12523]
SENSOR CONTROL | SENSORS_GET_VALUES | ~ | Two plain text lines are sent, not bracketed command format: line 1 starts 'ADC: ...'; line 2 starts 'I2C: ...' | [SENSORS_GET_VALUES]->[NOK][reason][TS] | Returns latest ADC and I2C converted values. Requires app RUNNING and all sensor-valid flags set. | SENSORS_GET_VALUES | ADC: ADCts:..., Iin:... / I2C: Pts:..., P:...
SENSOR CONTROL | SENSORS_GET_STATE | ~ | [SENSORS_GET_STATE]->[ADC_STATE][I2C_STATE][TS] |  | Returns ADC and I2C acquisition state strings. | SENSORS_GET_STATE | [SENSORS_GET_STATE]->[ADC_ACQ_STATE_DONE][I2C_ACQ_STATE_IDLE][12523]
SENSOR CONTROL | SENSORS_GET_LAST_ERROR | ~ | [SENSORS_GET_LAST_ERROR]->[ADC error string][ADC error code][I2C error string][I2C error code][TS] |  | Returns last ADC and I2C acquisition errors. | SENSORS_GET_LAST_ERROR | [SENSORS_GET_LAST_ERROR]->[No error][ADC_ACQ_ERROR_NONE][No error][I2C_ACQ_ERROR_NONE][12523]
SENSOR CONTROL | SENSORS_CLEAR_ERROR | ~ | [SENSORS_CLEAR_ERROR]->[OK][TS] |  | Clears ADC and I2C acquisition errors. | SENSORS_CLEAR_ERROR | [SENSORS_CLEAR_ERROR]->[OK][12523]
APP CONTROL | APP_PING | ~ | [APP_PING]->[OK][TS] |  | Heartbeat / presence check. | APP_PING | [APP_PING]->[OK][12523]
APP CONTROL | APP_GET_COMMANDS | ~ | Header: [APP_GET_COMMANDS]->[Count,N][TS]; then N lines [APP_GET_COMMANDS]->[CMD]; then [APP_GET_COMMANDS]->[END][TS] |  | Returns command list from g_cmdTable. | APP_GET_COMMANDS | [APP_GET_COMMANDS]->[Count,70][12523]
APP CONTROL | APP_GET_LIMITS | ~ | Header: [APP_GET_LIMITS]->[Count,18][TS]; then 18 text lines with current limits; then [APP_GET_LIMITS]->[END][TS] | [APP_GET_LIMITS]->[NOK][Cannot get limits while process is running][TS] | Returns current runtime limits and intervals. Some values are dynamic and depend on current parameter set. | APP_GET_LIMITS | [APP_GET_LIMITS]->[Count,18][12523]
APP CONTROL | APP_RESET | ~ | [APP_RESET]->[OK][TS] |  | Replies OK, waits 10 ms, then triggers MCU reset. | APP_RESET | [APP_RESET]->[OK][12523]
APP CONTROL | APP_GET_PROC_TIME | ~ | [APP_GET_PROC_TIME]->[OK][TS] |  | Implemented response currently only returns OK + timestamp via Comm_ReplyOk(). | APP_GET_PROC_TIME | [APP_GET_PROC_TIME]->[OK][12523] | Code/comment mismatch: handler does not return process uptime value.
APP CONTROL | APP_SET_VACUUM_EN | int/bool; >0 = on, <=0 = off | [APP_SET_VACUUM_EN][arg]->[OK][TS] | [APP_SET_VACUUM_EN]->[NOK][reason][TS] or [APP_SET_VACUUM_EN][arg]->[NOK][reason][TS] | Controls vacuum switch output. Blocked while process active or app not RUNNING. | APP_SET_VACUUM_EN 1 | [APP_SET_VACUUM_EN][1]->[OK][12523]
APP CONTROL | APP_GET_VACUUM_EN | ~ | [APP_GET_VACUUM_EN]->[1¦0][TS] |  | Reads vacuum switch state. | APP_GET_VACUUM_EN | [APP_GET_VACUUM_EN]->[1][12523]
APP CONTROL | APP_SET_CHECK_VACUUM | int/bool; >0 = enable, <=0 = disable | [APP_SET_CHECK_VACUUM][arg]->[OK][TS] | [APP_SET_CHECK_VACUUM]->[NOK][reason][TS] or [APP_SET_CHECK_VACUUM][arg]->[NOK][reason][TS] | Enables/disables vacuum safety check. Blocked while process active or app not RUNNING. | APP_SET_CHECK_VACUUM 1 | [APP_SET_CHECK_VACUUM][1]->[OK][12523]
APP CONTROL | APP_GET_CHECK_VACUUM | ~ | [APP_GET_CHECK_VACUUM]->[1¦0][TS] |  | Returns current vacuum-check enable flag. | APP_GET_CHECK_VACUUM | [APP_GET_CHECK_VACUUM]->[0][12523]
APP CONTROL | APP_DO_LASER_PWR_TEST | ~ | [APP_DO_LASER_PWR_TEST]->[OK][TS] | [APP_DO_LASER_PWR_TEST]->[NOK][reason][TS] | Starts calibration / laser power test. Requires app RUNNING, no active process, acceptable vacuum if check enabled, and Peltier enabled. | APP_DO_LASER_PWR_TEST | [APP_DO_LASER_PWR_TEST]->[OK][12523]
APP CONTROL | APP_GET_LASER_TEST_RESULT | ~ | [APP_GET_LASER_TEST_RESULT]->[OK][TS] | [APP_GET_LASER_TEST_RESULT]->[NOK][Laser power test failed or not completed][TS] | Returns OK only when calibration result flag is true. | APP_GET_LASER_TEST_RESULT | [APP_GET_LASER_TEST_RESULT]->[OK][12523]
APP CONTROL | APP_GET_LASER_TEST_DATA | ~ | Header: [APP_GET_LASER_TEST_DATA]->[Count,N][TS]; then N plain lines 'idx,T1,T2'; then [APP_GET_LASER_TEST_DATA]->[END][TS] | [APP_GET_LASER_TEST_DATA]->[NOK][reason][TS] | Returns captured calibration samples. Blocked while any process is active. | APP_GET_LASER_TEST_DATA | [APP_GET_LASER_TEST_DATA]->[Count,100][12523] | Data lines are NOT wrapped in [APP_GET_LASER_TEST_DATA]->[...] format.
APP CONTROL | APP_GET_FW_VERSION | ~ | [APP_GET_FW_VERSION]->[FW_VERSION][SW_GIT_HASH][TS] |  | Returns firmware version and software git hash from parameters. | APP_GET_FW_VERSION | [APP_GET_FW_VERSION]->[V15.1][abc1234][12523]
APP CONTROL | APP_GET_HW_VERSION | ~ | [APP_GET_HW_VERSION]->[HW_VERSION][HW_GIT_HASH][TS] |  | Returns hardware version and hardware git hash from parameters. | APP_GET_HW_VERSION | [APP_GET_HW_VERSION]->[V2.2][def5678][12523]
APP CONTROL | APP_GET_STATE | ~ | [APP_GET_STATE]->[APP_STATE_*][TS] |  | Returns current application state string. | APP_GET_STATE | [APP_GET_STATE]->[APP_STATE_RUNNING][12523]
APP CONTROL | APP_GET_LAST_ERROR | ~ | [APP_GET_LAST_ERROR]->[Error string][APP_ERROR_CODE][TS] |  | Returns application last error string and code. | APP_GET_LAST_ERROR | [APP_GET_LAST_ERROR]->[No error][APP_ERROR_NONE][12523]
APP CONTROL | APP_CLEAR_ERROR | ~ | [APP_CLEAR_ERROR]->[OK][TS] |  | Attempts to clear application error state. Only has effect if current app state is APP_STATE_ERROR; then app moves to STABILIZING. | APP_CLEAR_ERROR | [APP_CLEAR_ERROR]->[OK][12523]
APP CONTROL | APP_SET_GENERAL_PARAMS | VacMax_mbar,RhMin_dP,RhMax_dP,TBoardMin_dC,TBoardMax_dC,THeatsinkMin_dC,THeatsinkMax_dC | [APP_SET_GENERAL_PARAMS][args]->[OK][TS] | [APP_SET_GENERAL_PARAMS][args]->[NOK][reason][TS] | Sets general safety/operation limits. Blocked while process active or app not RUNNING. | APP_SET_GENERAL_PARAMS 700,0,500,100,600,100,600 | [APP_SET_GENERAL_PARAMS][700,0,500,100,600,100,600]->[OK][12523]
APP CONTROL | APP_SET_PELTIER_PARAMS | LaserTempMin_dC,LaserTempMax_dC,LaserTempNominal_dC,Kp,Ki,Kd,Div,iMax | [APP_SET_PELTIER_PARAMS][args]->[OK][TS] | [APP_SET_PELTIER_PARAMS][args]->[NOK][reason][TS] | Sets Peltier operating range and PID values. Blocked while process active or app not RUNNING. | APP_SET_PELTIER_PARAMS 200,300,250,10000,10,0,8,10000 | [APP_SET_PELTIER_PARAMS][200,300,250,10000,10,0,8,10000]->[OK][12523]
APP CONTROL | APP_SET_LASER_PARAMS | MaxFireDuration_ms,RedDotPowerPercent | [APP_SET_LASER_PARAMS][args]->[OK][TS] | [APP_SET_LASER_PARAMS][args]->[NOK][reason][TS] | Sets max fire duration and red-dot power. Blocked while process active or app not RUNNING. | APP_SET_LASER_PARAMS 100,25 | [APP_SET_LASER_PARAMS][100,25]->[OK][12523]
APP CONTROL | APP_SET_CALIB_PARAMS | TolerancePercent,CalibPulseDuration_ms | [APP_SET_CALIB_PARAMS][args]->[OK][TS] | [APP_SET_CALIB_PARAMS][args]->[NOK][reason][TS] | Sets calibration tolerances and pulse duration. Blocked while process active or app not RUNNING. | APP_SET_CALIB_PARAMS 10,100 | [APP_SET_CALIB_PARAMS][10,100]->[OK][12523]
APP CONTROL | APP_GET_GENERAL_PARAMS | ~ | [APP_GET_GENERAL_PARAMS]->[VacMax_mbar:...,RhMin_dP:...,RhMax_dP:...,TBoardMin_dC:...,TBoardMax_dC:...,THeatsinkMin_dC:...,THeatsinkMax_dC:...][TS] |  | Returns current general parameters. | APP_GET_GENERAL_PARAMS | [APP_GET_GENERAL_PARAMS]->[VacMax_mbar:700,RhMin_dP:0,RhMax_dP:500,TBoardMin_dC:100,TBoardMax_dC:600,THeatsinkMin_dC:100,THeatsinkMax_dC:600][12523]
APP CONTROL | APP_GET_PELTIER_PARAMS | ~ | [APP_GET_PELTIER_PARAMS]->[LaserTempMin_dC:...,LaserTempMax_dC:...,LaserTempNominal_dC:...,Kp:...,Ki:...,Kd:...,Div:...,iMax:...][TS] |  | Returns current Peltier parameters. | APP_GET_PELTIER_PARAMS | [APP_GET_PELTIER_PARAMS]->[LaserTempMin_dC:200,LaserTempMax_dC:300,LaserTempNominal_dC:250,Kp:10000,Ki:10,Kd:0,Div:8,iMax:10000][12523]
APP CONTROL | APP_GET_LASER_PARAMS | ~ | [APP_GET_LASER_PARAMS]->[MaxFireDuration_ms:...,RedDotPowerPercent:...][TS] |  | Returns current laser parameters. | APP_GET_LASER_PARAMS | [APP_GET_LASER_PARAMS]->[MaxFireDuration_ms:100,RedDotPowerPercent:25][12523]
APP CONTROL | APP_GET_CALIB_PARAMS | ~ | Intended format string is [APP_GET_CALIB_PARAMS]->[TolerancePercent,CalibPulseDuration_ms][TS] but current code omits value placeholders. |  | Handler tries to print calibration parameters, but format string bug means values are not inserted as written. | APP_GET_CALIB_PARAMS | [APP_GET_CALIB_PARAMS]->[TolerancePercent,CalibPulseDuration_ms][12523] | Bug in implementation: values are passed to Comm_PrintUSB but format string contains no % placeholders for them.
APP CONTROL | APP_BOOT | ~ |  | [APP_BOOT]->[NOK][Boot command is not implemented.][TS] | Currently not implemented. | APP_BOOT | [APP_BOOT]->[NOK][Boot command is not implemented.][12523]
APP CONTROL | APP_STORE_PARAMS_TO_NVM | ~ | [APP_STORE_PARAMS_TO_NVM]->[OK][TS] | [APP_STORE_PARAMS_TO_NVM]->[NOK][reason][TS] | Stores current parameters to data EEPROM. Blocked while process active or app not RUNNING. | APP_STORE_PARAMS_TO_NVM | [APP_STORE_PARAMS_TO_NVM]->[OK][12523]
APP CONTROL | APP_STORE_CALIB_DATA | ~ | [APP_STORE_CALIB_DATA]->[OK][TS] | [APP_STORE_CALIB_DATA]->[NOK][reason][TS] | Copies captured calib data into reference array (average of T1/T2) and marks refDataValid=true. Does not itself write NVM; use APP_STORE_PARAMS_TO_NVM afterwards. | APP_STORE_CALIB_DATA | [APP_STORE_CALIB_DATA]->[OK][12523]
APP CONTROL | APP_SET_PERIODIC_REPORT_EN | int/bool; >0 = enable, <=0 = disable | [APP_SET_PERIODIC_REPORT_EN][arg]->[OK][TS] | [APP_SET_PERIODIC_REPORT_EN][arg]->[NOK][Invalid argument(s)][TS] | Enables/disables periodic report messages. | APP_SET_PERIODIC_REPORT_EN 1 | [APP_SET_PERIODIC_REPORT_EN][1]->[OK][12523]
APP CONTROL | APP_GET_PERIODIC_REPORT_EN | ~ | [APP_GET_PERIODIC_REPORT_EN]->[1¦0][TS] |  | Returns periodic report enable flag. | APP_GET_PERIODIC_REPORT_EN | [APP_GET_PERIODIC_REPORT_EN]->[0][12523]
APP CONTROL | APP_SET_PELTIER_DEFAULT_EN | int/bool; >0 = enable startup Peltier, <=0 = disable | [APP_SET_PELTIER_DEFAULT_EN][arg]->[OK][TS] | [APP_SET_PELTIER_DEFAULT_EN][arg]->[NOK][Invalid argument(s)][TS] | Sets whether Peltier is automatically enabled after stabilization on startup. | APP_SET_PELTIER_DEFAULT_EN 1 | [APP_SET_PELTIER_DEFAULT_EN][1]->[OK][12523]
APP CONTROL | APP_GET_PELTIER_DEFAULT_EN | ~ | [APP_GET_PELTIER_DEFAULT_EN]->[1¦0][TS] |  | Returns startup default for Peltier enable. | APP_GET_PELTIER_DEFAULT_EN | [APP_GET_PELTIER_DEFAULT_EN]->[1][12523]
APP CONTROL | APP_RESET_TO_DEFAULTS | ~ | [APP_RESET_TO_DEFAULTS]->[OK][TS] | [APP_RESET_TO_DEFAULTS]->[NOK][reason][TS] | Resets in-memory parameters to compiled defaults. Blocked while process active or app not RUNNING. | APP_RESET_TO_DEFAULTS | [APP_RESET_TO_DEFAULTS]->[OK][12523]
Note on A1: Derived from the app.c code snippet provided by the user in this chat on 2026-04-17.

## Async Messages
Message | Payload / Format | When sent | Notes | Example
SYSTEM_STABILIZED | [SYSTEM_STABILIZED]->[TS] | After APP_STATE_STABILIZING has lasted APP_STABILIZATION_TIME_MS (1000 ms), just before switching to APP_STATE_RUNNING. | Not in command table, but emitted from App_Task(). | [SYSTEM_STABILIZED]->[12523]
APP_ERROR_HAPPENED | [APP_ERROR_HAPPENED]->[Error string][APP_ERROR_CODE][TS] | When g_appCtx.sendReportOnError is set after App_Fail(). | This is the actual code format; old draft had field order reversed. | [APP_ERROR_HAPPENED]->[Pressure above safety limit][APP_ERROR_UNSAFE_CONDITIONS][12523]
APP_WATCHDOG_TRIGGERED | [APP_WATCHDOG_TRIGGERED]->[TS] | When g_appCtx.sendReportOnWatchdogTimeout is set. | Current implementation sends only timestamp. There is also an older duplicate spec in the draft with error text, but not in this code. | [APP_WATCHDOG_TRIGGERED]->[12523]
APP_LASER_PWR_TEST_FINISHED | [APP_LASER_PWR_TEST_FINISHED]->[OK¦NOK][TS] | When calibration/power test completes and OnCalibComplete() sets report flag. |  | [APP_LASER_PWR_TEST_FINISHED]->[OK][12523]
APP_TARGET_SEQ_FINISHED | [APP_TARGET_SEQ_FINISHED]->[OK][TARGET_STATE_*][TS] | When target sequence completes and OnTargetSequenceCompleted() sets report flag. | Current implementation always prints OK plus final target state string. | [APP_TARGET_SEQ_FINISHED]->[OK][TARGET_STATE_COMPLETED][12523]
APP_LOAD_PARAMS_FAILED | [APP_LOAD_PARAMS_FAILED]->[Error string][APP_ERROR_CODE][TS] | Once after startup if parameters could not be loaded from NVM and defaults are used. |  | [APP_LOAD_PARAMS_FAILED]->[Failed to load parameters from NVM, using defaults.][APP_ERROR_NVM][12523]
Periodic report: ADC | ADC: ADCts:<ts>, Iin:<mA>, I660:<mA>, I808:<mA>, I980:<mA>, I1064:<mA>, Pwr:<mV>, Tlaser:<dC>, Vpelt:<mV>, Theat:<dC>, Tmos:<dC>, T1:<dC>, T2:<dC>, Gx:<raw>, Gy:<raw> | Every APP_REPORT_PERIOD_MS (5000 ms) when APP_SET_PERIODIC_REPORT_EN is enabled and app is RUNNING. | Plain text line, not wrapped in [CMD]->[...] format.
Periodic report: I2C | I2C:Pts:<ts>, SHTts:<ts>, P:<mBar>, PT:<dC>, Tamb:<dC>, RH:<dP> | Every APP_REPORT_PERIOD_MS (5000 ms) when periodic reporting is enabled. | Plain text line, not wrapped in [CMD]->[...] format.
Periodic report: APP | APP: <APP_STATE>\nADC:<ADC_STATE>\nI2C:<I2C_STATE>\nPELTIER:<PELTIER_STATE>\nLASER:<LASER_STATE>\nTARGET:<TARGET_STATE>\nCALIB:<CALIB_STATE>\nScheduler<counter> | Every APP_REPORT_PERIOD_MS (5000 ms) when periodic reporting is enabled. | Contains embedded newline characters inside one transmitted string.
Note on A1: Derived from the app.c code snippet provided by the user in this chat on 2026-04-17.

## Defaults
Group | Macro / Parameter | Default value | Units / Type
General | APP_DEFAULT_LASER_MAX_FIRE_DURATION_MS | 100 | ms
General | APP_DEFAULT_RED_DOT_POWER_PERCENT | 25 | %
Peltier | APP_DEFAULT_LASER_TEMP_MIN_dC | 200 | d°C
Peltier | APP_DEFAULT_LASER_TEMP_MAX_dC | 300 | d°C
Peltier | APP_DEFAULT_LASER_TEMP_NOMINAL_dC | 250 | d°C
Peltier | APP_DEFAULT_PELTIER_PID_KP | 10000
Peltier | APP_DEFAULT_PELTIER_PID_KI | 10
Peltier | APP_DEFAULT_PELTIER_PID_KD | 0
Peltier | APP_DEFAULT_PELTIER_PID_DIV | 8
Peltier | APP_DEFAULT_PELTIER_PID_IMAX | 10000
Calib | APP_DEFAULT_CALIB_PULSE_DURATION_MS | 100 | ms
Calib | APP_DEFAULT_CALIB_TOLERANCE_PERCENT | 10 | %
General | APP_DEFAULT_PERIODIC_REPORT_EN | FALSE | bool
General | APP_DEFAULT_PELTIER_DEFAULT_EN | TRUE | bool
Safety | APP_DEFAULT_VACUUM_MAX_mbar | 700 | mbar
Safety | APP_DEFAULT_RH_MIN_PERCENT_dP | 0 | d%
Safety | APP_DEFAULT_RH_MAX_PERCENT_dP | 500 | d%
Safety | APP_DEFAULT_TBOARD_MAX_TEMP_dC | 600 | d°C
Safety | APP_DEFAULT_TBOARD_MIN_TEMP_dC | 100 | d°C
Safety | APP_DEFAULT_HEATSINK_TEMP_MAX_dC | 600 | d°C
Safety | APP_DEFAULT_HEATSINK_TEMP_MIN_dC | 100 | d°C
Safety | APP_DEFAULT_CHECK_VACUUM | FALSE | bool
Timing | APP_STABILIZATION_TIME_MS | 1000 | ms
Timing | APP_REPORT_PERIOD_MS | 5000 | ms
Timing | APP_PRESSURE_PERIOD_MS | 10 | ms
Timing | APP_SHT40_PERIOD_MS | 500 | ms
Timing | APP_PELTIER_PERIOD_MS | 10 | ms
Timing | APP_BASE_TICK_MS | 2 | ms
Note on A1: Derived from the app.c code snippet provided by the user in this chat on 2026-04-17.

## Notes
General response patterns
Case | Format
Generic OK, no args | [CMD]->[OK][TS]
Generic OK, with args echoed | [CMD][arg_string]->[OK][TS]
Generic NOK, no args echoed | [CMD]->[NOK][reason][TS]
Generic NOK, with args echoed | [CMD][arg_string]->[NOK][reason][TS]
Unknown command | [UNKNOWN]->[NOK][Unknown CMD][TS]
Command buffer overflow | ERR: Command too long
Command list multi-line response | [CMD]->[Count,N][TS] ... [CMD]->[END][TS]
Key implementation notes / discrepancies vs older draft
Topic | Detail
TARGET_SET_MODE expects integer enum, not text | The handler parses the argument with Comm_ParseLong() and casts to TargetMode_t. Old sheet text 'SINGLE/AUTO' is no longer correct.
APP_BOOT is stubbed | Command always returns [APP_BOOT]->[NOK][Boot command is not implemented.][TS].
APP_GET_PROC_TIME does not return uptime | Current implementation simply calls Comm_ReplyOk(cmd), so result is [APP_GET_PROC_TIME]->[OK][TS].
APP_GET_CALIB_PARAMS format string bug | The handler passes tolerance and pulse-duration values, but the format string contains no % placeholders for them.
LASER_SET_CHANNEL_PWR currently rejects 0 | Validation requires each channel power to be 1..100. In contrast TARGET_SET_NEW_TARGET allows per-channel power 0..100.
Several outputs are plain text, not [CMD]->[...] formatted | SENSORS_GET_VALUES, APP_GET_LASER_TEST_DATA sample lines, and periodic reports use raw text lines.
APP_TARGET_SEQ_FINISHED payload differs from old draft | Actual code sends [APP_TARGET_SEQ_FINISHED]->[OK][<target state string>][TS].
APP_ERROR_HAPPENED field order differs from old draft | Actual code sends [APP_ERROR_HAPPENED]->[Error string][APP_ERROR_CODE][TS].
APP_WATCHDOG_TRIGGERED currently only includes timestamp | No error code or message is included in this implementation.
APP_STORE_CALIB_DATA does not write flash by itself | It only copies averaged T1/T2 data into params.calibParams.calibDataRef and sets refDataValid=true. Persist with APP_STORE_PARAMS_TO_NVM.
Note on A1: Derived from the app.c code snippet provided by the user in this chat on 2026-04-17.