## Commands
CMD | Parameters | Return | Description | Example
LASER CONTROL
LASER_SET_ARM_EN | [bool (0-1)] | [LASER_SET_ARM_EN]->[OK/NOK] | Set laser armed state. | LASER_SET_ARM_EN 1/0
LASER_GET_ARM_EN | [~] | [LASER_GET_ARM_EN]->[1/0] | Returns if laser is armed. | LASER_GET_ARM_EN
LASER_SET_RED_DOT_EN | [bool (0-1)] | [LASER_SET_RED_DOT_EN]->[OK/NOK] | Activates 660nm red laser with 25% power for pointer purposes. | LASER_SET_RED_DOT_EN 1/0
LASER_GET_RED_DOT_EN | [~] | [LASER_GET_RED_DOT_EN]->[1/0] | Returns if laser red dot mode is active. | LASER_GET_RED_DOT_EN
LASER_SET_CHANNEL_PWR | [808PWR (0-100), 980PWR (0-100), 1064PWR (0-100)] | [LASER_SET_CHANNEL_PWR]->[OK/NOK] | Sets laser channel powers in %. | LASER_SET_CHANNEL_PWR 20,25,50
LASER_GET_CHANNEL_PWR | [~] | [LASER_GET_CHANNEL_PWR]->[20,25,50] | Returns the set channel powers in %. | LASER_GET_CHANNEL_PWR
LASER_FIRE | [duration_ms (10-1000)] | [LASER_FIRE]->[OK/NOK] | Fires laser for specified duration in [ms]. | LASER_FIRE 20
LASER_STOP | [~] | [LASER_STOP]->[OK/NOK] | Stops laser firing immediately. | LASER_STOP
LASER_IS_ACTIVE | [~] | [LASER_IS_ACTIVE]->[1/0] | Returns if laser is currently firing. | LASER_IS_ACTIVE
LASER_GET_STATE | [~] | [LASER_GET_STATE]->[LASER internal state] | Returns laser controller internal state. LASER_STATE_DISARMED, LASER_STATE_ARMED_IDLE, LASER_STATE_FIRING, LASER_STATE_ERROR | LASER_GET_STATE
LASER_GET_LAST_ERROR | [~] | [LASER_GET_LAST_ERROR]->[ERROR_CODE],[Error string],[Timestamp] | Requests the last occurred error string. | LASER_GET_LAST_ERROR
LASER_CLEAR_ERROR | [~] | [LASER_CLEAR_ERROR]->[OK/NOK] | Clears last error, resets controller state, disarms laser. | LASER_CLEAR_ERROR
PELTIER CONTROL
PELTIER_SET_COOLING_EN | [bool (0-1)] | [PELTIER_SET_COOLING_EN]->[OK/NOK] | Sets the laser temperature stabilization on/off. | PELTIER_SET_COOLING_EN 1/0
PELTIER_GET_COOLING_EN | [~] | [PELTIER_GET_COOLING_EN]->[1/0] | Returns if laser cooling is enabled. | PELTIER_GET_COOLING_EN
PELTIER_GET_STATE | [~] | [PELTIER_GET_STATE]->[PELTIER internal state] | Returns peltier controller internal state. PELTIER_STATE_DISABLED, PELTIER_STATE_ENABLED, PELTIER_STATE_ERROR | PELTIER_GET_STATE
PELTIER_GET_LAST_ERROR | [~] | [PELTIER_GET_LAST_ERROR]->[ERROR_CODE],[Error string],[Timestamp] | Requests the last occurred error string. | PELTIER_GET_LAST_ERROR
PELTIER_CLEAR_ERROR | [~] | [PELTIER_CLEAR_ERROR]->[OK/NOK] | Clears last error, resets controller state, disables cooling. | PELTIER_CLEAR_ERROR
TARGET CONTROL
TARGET_SET_POS | [X (0-4095),Y (0-4095)] | [TARGET_SET_POS]->[OK/NOK] | Sets the galvo to the given position. | TARGET_SET_POS 1500,2000
TARGET_GET_POS | [~] | [TARGET_GET_POS]->[X,Y] | Returns the last set position. This is not the galvo position feedback signal. | TARGET_GET_POS
TARGET_GET_FB_POS | [~] | [TARGET_GET_FB_POS]->[aX,bY] | Returns the galvo feedback position signal value. This is linearly related to the actual position, but not the same. | TARGET_GET_FB_POS
TARGET_SET_NEW_TARGET | [X0 (0-4095), Y0 (0-4095), X1 (0-4095), Y1 (0-4095), 808PWR (0-100), 980PWR (0-100), 1064PWR (0-100), duration_ms (10-1000)] | [TARGET_SET_NEW_TARGET]->[OK/NOK] | Adds a new target point or line to the target list. If target is a single point, set X0=X1, Y0=Y1. Max. number of targets is 256. | TARGET_SET_NEW_TARGET 1500,1000,2000,1500,20,10,15,50
TARGET_CLEAR_TARGETS | [~] | [TARGET_CLEAR_TARGETS]->[OK/NOK] | Clears target list. | TARGET_CLEAR_TARGETS
TARGET_GET_TARGET | [ID (0-255)] | [TARGET_GET_TARGET]->[5,1500,1000,2000,1500,20,10,15,50] | Returns the given target at the specified ID. If no target at the given ID, NOK is returned. | TARGET_GET_TARGET 5
TARGET_GET_ALL_TARGETS | [~] | [TARGET_GET_ALL_TARGETS]->[Count,3] [TARGET_GET_ALL_TARGETS]->[0,1500,1000,2000,1500,20,10,15,50] [TARGET_GET_ALL_TARGETS]->[1,2000,1500,2200,1300,20,15,20,40] [TARGET_GET_ALL_TARGETS]->[2,1600,1400,1700,1700,25,30,25,30][TARGET_GET_ALL_TARGETS]->END | Returns all valid targets. The first parameter is the target ID. | TARGET_GET_ALL_TARGETS
TARGET_SET_MODE | [SINGLE/AUTO] | [TARGET_SET_MODE]->[OK/NOK] | Sets the target acquisition mode to single (manual re-trigger) or auto. | TARGET_SET_MODE SINGLE
TARGET_START | [~] | [TARGET_START]->[OK/NOK] | Starts shooting at targets. | TARGET_START
TARGET_HALT | [~] | [TARGET_HALT]->[OK/NOK] | Halts shooting at targets. Can be continued by calling TARGET_CONTINUE. | TARGET_HALT
TARGET_STOP | [~] | [TARGET_STOP]->[OK/NOK] | Stops shooting at targets and resets sequence. Calling START again re-runs the sequence. Cannot be continued. | TARGET_STOP
TARGET_CONTINUE | [~] | [TARGET_CONTINUE]->[OK/NOK] | Continues halted sequence. | TARGET_CONTINUE
TARGET_GET_STATE | [~] | [TARGET_GET_STATE]->[TARGET internal state] | Returns target controller internal state. TARGET_SEQ_STATUS_IDLE, TARGET_SEQ_STATUS_RUNNING, TARGET_SEQ_STATUS_HALTED, TARGET_SEQ_STATUS_COMPLETED, TARGET_SEQ_STATUS_ERROR | TARGET_GET_STATE
TARGET_GET_LAST_ERROR | [~] | [TARGET_GET_LAST_ERROR]->[ERROR_CODE],[Error string],[Timestamp] | Requests the last occurred error string. | TARGET_GET_LAST_ERROR
TARGET_CLEAR_ERROR | [~] | [TARGET_CLEAR_ERROR]->[OK/NOK] | Clears last error, resets controller state. | TARGET_CLEAR_ERROR
SENSOR CONTROL
SENSORS_GET_VALUES | [~] | [SENSORS_GET_VALUES]->[All sensor values as a list. *Listed in sheet "SensorValues".] | Returns all analog and digital sensors latest values as a list. | SENSORS_GET_VALUES
SENSORS_GET_STATE | [~] | [SENSORS_GET_STATE]->[SENSORS internal state] | Returns target controller internal state. ADC_ACQ_STATE_IDLE, ADC_ACQ_STATE_START_GROUP, ADC_ACQ_STATE_READING, ADC_ACQ_STATE_DONE, ADC_ACQ_STATE_ERROR | SENSORS_GET_STATE
SENSORS_GET_LAST_ERROR | [~] | [SENSORS_GET_LAST_ERROR]->[ERROR_CODE],[Error string],[Timestamp] | Requests the last occurred error string. | SENSORS_GET_LAST_ERROR
SENSORS_CLEAR_ERROR | [~] | [SENSORS_CLEAR_ERROR]->[OK/NOK] | Clears last error, resets controller state. | SENSORS_CLEAR_ERROR
APP CONTROL
APP_PING | [~] | [APP_PING]->[OK] | Can be used for heartbeat/presence detection. | APP_PING
APP_GET_COMMANDS | [~] | [APP_GET_COMMANDS]->[Command list.] | Returns the content of this file. | APP_GET_COMMANDS
APP_GET_LIMITS | [~] | [APP_GET_LIMITS]->[String with parameter limits] | Returns the parameter limits. | APP_GET_LIMITS
APP_RESET | [~] | [APP_RESET]->[OK] | Initiates a SW reset to the microcontroller. | APP_RESET
APP_GET_PROC_TIME | [~] | [APP_GET_PROC_TIME]->[Process time since powerup in ms.] | Returns the internal systick counter value. Can be used to check freshness of responses. | APP_GET_PROC_TIME
APP_DO_LASER_PWR_TEST | [~] | [APP_DO_LASER_PWR_TEST]->[OK/NOK] | Performs laser power test on the test target. | APP_DO_LASER_TEST
APP_GET_LASER_TEST_RESULT | [~] | [APP_GET_LASER_TEST_RESULT]-[PASS/FAIL/BUSY] | Returns if laser power test was successful. | APP_GET_LASER_TEST_RESULT
APP_GET_LASER_TEST_DATA | [~] | [APP_GET_LASER_TEST_DATA]->[Measured heating curve. *Listed in sheet "PowerTestData".] | Returns the sampled temperature points during the test. | APP_GET_LASER_TEST_DATA
APP_GET_FW_VERSION | [~] | [APP_GET_FW_VERSION]->[V15.1,2026.03.06.,GIT_HASH] | Returns active FW version string and git commit hash. | APP_GET_FW_VERSION
APP_GET_HW_VERSION | [~] | [APP_GET_HW_VERSION]->[V2.2,2025.09.26.,GIT_HASH] | Returns current HW version string and git commit hash. | APP_GET_HW_VERSION
APP_GET_STATE | [~] | [APP_GET_STATE]->[Application internal state] | Returns the overall state of the application. APP_STATE_IDLE, APP_STATE_ERROR, APP_STATE_HARD_FAULT | APP_GET_STATE
APP_GET_LAST_ERROR | [~] | [APP_GET_LAST_ERROR]->[ERROR_CODE],[Error string],[Timestamp] | Requests the last occurred error string. | APP_GET_LAST_ERROR
APP_CLEAR_ERROR | [~] | [APP_CLEAR_ERROR]->[OK/NOK] | Clears last error, resets every state. | APP_CLEAR_ERROR
ASYNC MESSAGES |  |  | Sent by the device if something unexpected happens or some event finishes.
APP_ERROR_HAPPENED | [~] | [APP_ERROR_HAPPENED]->[ERROR_CODE],[Error string],[Timestamp] | Error happened at some point. Call APP_GET_LAST_ERROR to request error message, or APP_CLEAR_ERROR to try recovery from error. | [APP_ERROR_HAPPENED]->Laser temp out of regulation range. Tlaser:32.2°C
APP_HARD_FAULT_HAPPENED | [~] | [APP_HARD_FAULT_HAPPENED]->[ERROR_CODE],[Error string],[Timestamp] | Hard fault happened, no recovery possible. Call APP_GET_LAST_ERROR, or try reset device. | [APP_HARD_FAULT_HAPPENED]->Pressure sensor unavailable. I2C DMA error.
TARGET_SEQ_FINISHED | [~] | [TARGET_SEQ_FINISHED]->[OK/NOK],[Timestamp] | Target sequence finished. If successful returns OK, otherwise NOK. | [TARGET_SEQ_FINISHED]->OK
APP_WATCHDOG_TRIGGERED | [~] | [APP_WATCHDOG_TRIGGERED]->[ERROR_CODE],[Error string],[Timestamp] | Watchdog reset was triggered. Device starts from reset state. | [APP_WATCHDOG_TRIGGERED]->Application deadlock.
APP_LASER_PWR_TEST_FINISHED | [~] | [APP_LASER_PWR_TEST_FINISHED]->[OK/NOK],[Timestamp] | Laser power test finished. Returns OK in case of success, NOK otherwise. Obtain results by calling APP_GET_LASER_TEST_DATA if needed. | [APP_LASER_PWR_TEST_FINISHED]->OK
NOK HANDLING
If a command returns NOK, the return string is formatted as follows: | [CMD]->[NOK],[Parameters sent], [ERROR_CODE], [Error string], [Timestamp] | eg.: [SET_LASER_ARM_EN]->[NOK],[1],[ERR_NO_VACUUM],[No vacuum present. P: 967mBar],[26342] | Error codes are listed on sheet "ErrorCodes"
TARGETING EXAMPLES
Single target acquisition:
| LASER_SET_RED_DOT_EN 1 | Optional, LASER_SET_ARM_EN 1 must be called if not already armed.
| LASER_SET_CHANNEL_PWR 10,10,0 | 0 means that channel is inactive, no separate set-active-channels command.
| TARGET_SET_POS 1700,2000 | Move galvo to given spot.
| LASER_FIRE 50 | Fires laser for 50 ms.
Sequence acquisition (AUTO):
| TARGET_SET_MODE AUTO | Sets automatic mode. Targets will be captured without waiting for user trigger.
| TARGET_CLEAR_TARGETS | Clear all targets.
| TARGET_SET_NEW_TARGET 1500,1000,1500,1000,20,10,15,50 | Add new target point.
| TARGET_SET_NEW_TARGET 1500,1000,2000,1500,20,10,15,50 | Add new target line.
| TARGET_SET_NEW_TARGET 1500,1000,2000,1500,20,10,15,50 | Add new target line.
| TARGET_SET_NEW_TARGET 1500,1000,1500,1000,20,10,15,50 | Add new point.
| TARGET_START | Starts target acquisition and firing.

## SensorValues
Sensor value list | Comment
[SENSORS_GET_VALUES]->[Count,19]
INPUT_CURR [mA]: inputCurrent_mA
LASER660_CURR [mA]: laser660Curr_mA
LASER808_CURR [mA]: laser808Curr_mA
LASER980_CURR [mA]: laser980Curr_mA
LASER1064_CURR [mA]: laser1064Curr_mA
LASER_POWER [mV]: laserPower_mV
LASER_TEMP [C]: laserTemp_C
PELTIER_VOLTAGE [mV]: peltierVoltage_mV
HEATSINK_TEMP [C]: heatsinkTemp_C
MOSFET_TEMP [C]: mosfetTemp_C
TARGET1_TEMP [C]: target1Temp_C
TARGET2_TEMP [C]: target2Temp_C
GALVO_POS_X [raw]: galvoPosX_raw
GALVO_POS_Y [raw]: galvoPosY_raw
LASER660_OFFSET [mA]: laser660Curr_mA
LASER808_OFFSET [mA]: laser808Curr_mA
LASER980_OFFSET [mA]: laser980Curr_mA
LASER1064_OFFSET [mA]: laser1064Curr_mA
UPDATE_TIMESTAMP [ms]: updateTimestamp_ms
[SENSORS_GET_VALUES]->[END] | Multi-line message termination.

## PowerTestData
Laser power test data | Comment
time [ms],temp [°C]
[APP_GET_LASER_TEST_DATA]->[Count,256]
0,25.0
10,26.7
20,27.2
…
[APP_GET_LASER_TEST_DATA]->[END] | Multi-line message termination.

## ErrorCodes
Error code | Description
ERR_NO_VACUUM | No vacuum present. Eg. by calling LASER_SET_ARM_EN 1
ERR_LASER_TEMP | Laser temp out of normal range.
ERR_SEQ_IN_PROG | Sequence already running.
ERR_LASER_ACTIVE | Laser currently active. (LASER_FIRE was just sent)
ERR_GALVO_FAULT | Galvo fault pin IT occoured.
ERR_OUT_OF_RANGE | Parameter specified out of range.
ERR_SENS_FAILURE | Sensor failure occoured.
ERR_NOT_ARMED | Laser is not armed.
ERR_INVALID_CMD | Invalid command or missing parameters.
ERR_INVALID_OPERATION | Invalid operation in current state.
ERR_ENV_COND | Environmental conditions not met. Tboard and RHboard not in specified range.
…